3803 COMMUNICATION PROTOCOL

The serial data is output two times from SDO pin on each A/D conversion cycle. The data format satisfy JIS 7BIT transmission code and the baud rate is 2400, it means the receiver terminal can use RS-232 interface to read data.

### Single package

#### 0V

| -3V | ## 0 |  |  | 1 |  |
| --- | --- | --- | --- | --- | --- |
|  |  | LSB | MSB |  |  |

A single package include a start bit(0),D0~D6(7bit), a parity check bit(odd) and a stop bit(1). All the data stream is consisit of 110 x 2 bits.

all   package

|  | 0 |  | 1 | 0 |  | 1 | 0 |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | range |  |  | digit 3 |  |  | digit 2 |  |

|  | 0 |  | 1 | 0 |  | 1 | 0 |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | digit 1 |  |  | digit 0 |  |  | function |  |

|  | 0 |  | 1 | 0 |  | 1 | 0 |  | 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | status |  |  | option 1 |  |  | option 2 |  |

|  | 0 |  | 1 | 0 |  | 1 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | CR |  |  | LF |  |  |  |  |  |

Digit 0~3 data is exactly the receiver display number, but the status, function, range, option1, option2 data all need written program to judge what the data mean.

Note: LSB bit is sent first, then MSB bit.

The output of RS232 is always the current measured value. Once auto-power-off happen, the final data will not output.

RANGE :

| ##### Range | V | mA | uA | Ohm | Frequency |
| --- | --- | --- | --- | --- | --- |
| 0110000 | 400mV | 40mA | 400uA | 400 Ohm | ###### 4K HZ |
| 0110001 | 4V | 400mA | 4000uA | 4K Ohm | 40K HZ |
| 0110010 | 40V |  |  | 40K Ohm | 400K HZ |
| 0110011 | 400V |  |  | 400K Ohm | 4M HZ |
| 0110100 | 4000V |  |  | 4M Ohm | 40M HZ |
| 0110101 |  |  |  | 40M Ohm |  |

Because Continuity, Diode, and Current(A) measurement range are all fixed, so they are all the first code 0110000.

DIGIT0 ~3 :

| Digit | Display data |
| --- | --- |
| 0110000 | 0 |
| 0110001 | 1 |
| 0110010 | 2 |
| 0110011 | 3 |
| 0110100 | 4 |
| 0110101 | 5 |
| 0110110 | 6 |
| 0110111 | 7 |
| 0111000 | 8 |
| 0111001 | 9 |

If the input value is over-range(OL), the reading data DIGIT3=4, DIGIT2=0, DIGIT1=0, DIGIT 0=0, but under the frequency mode, DIGIT0~3 output the measured value.

FUNCTION : The definition code is not the same as FC1 ~ FC4.

| Digit | Measurement mode |
| --- | --- |
| 0111111 | A |
| 0111110 | uA |
| 0111101 | mA |
| 0111100 | Voltage |
| 0110111 |  |
| 0110110 | Continuity |
| 0111011 | Diode |
| 0111010 | Frequency |

STATUS :

| 0 | 1 | 1 | — | BATT | 0 | OL |
| --- | --- | --- | --- | --- | --- | --- |
| BIT7 | BIT6 | BIT5 | BIT4 | BIT3 | BIT2 | BIT1 |

— 	: “1” is for the negative (-).

BATT	: “1” is for the battery voltage is less than 2.3V ± 0.2V

OL	: “1” is for the input voltage over-range.

OPTION 1 :

| 0 | 1 | 1 | 0 | 0 | 0 | 0 |
| --- | --- | --- | --- | --- | --- | --- |
| BIT7 | BIT6 | BIT5 | BIT4 | BIT3 | BIT2 | BIT1 |

OPTION 2 :

| 0 | 1 | 1 | DC/AC | AUTO | MAN | APO |
| --- | --- | --- | --- | --- | --- | --- |
| BIT7 | BIT6 | BIT5 | BIT4 | BIT3 | BIT2 | BIT1 |

DC/AC	: “1” is for the DC mode and , Diode, Continuity, ADP, Frequency Function, “0” is for the AC mode

AUTO	: “1” is for auto range.

Man		: “1” is for the manual range.

APO		: “1” is for the auto-power-off function enabled.

CR(BACK)			: Transmission code “0001101”.

LF(LINE FEED)	: Transmission code “0001010”.

3804/3805 COMMUNICATION PROTOCOL

A. RS-232 protocol

- Baud Rate = 4800

- Data Bits = 6

- Parity = None

- Stop Bits = 1

B. Format Definition (11 bytes)

| Byte | Definition |
| --- | --- |
|  | Start Code |
|  | Function (0~16, refer to the definition for function) |
|  | Range |
|  | Polarity( 6=positive +, 9=Negaitive - ) |
|  | Value Digital 5 (if = 10, the Primary Value = OL) |
|  | Value Digital 4 |
|  | Value Digital 3 |
|  | Value Digital 2 |
|  | Value Digital 1 |
|  | Bit0: Low battery sign (1= low battery, 0= normal) Bit1: Auto-range( 1= manual, 0= auto) Bit2: % of mA display (1=enable, 0= disable) Bit3: Secondary Display (1=enable, 0=disable ) |
|  | Stop Code |

C. Command and Data codes Definition

| Definition | Stop bit | B5 | B4 | B3 | B2 | B1 | B0 | Start bit |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Start Code | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 |
| Stop Code | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 |
| Value 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |
| Value 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 |
| Value 2 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 |
| Value 3 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 |
| Value 4 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 |
| Value 5 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 |
| Value 6 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 |
| Value 7 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 |
| Value 8 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 |
| Value 9 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 |
| OVERLOAD(OL) | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 |

D. Function Definition

| Function | Stop | B5 | B4 | B3 | B2 | B1 | B0 | Start | Bit2 of byte10 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| DC V | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  |
| DC µA | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 |  |
| DC mA | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 |  |
| DC A | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 |  |
| Ts | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 |  |
| DIODE | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 |  |
| OHM | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 |  |
| Cx | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 |  |
| AC V | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 |  |
| AC µA | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 |  |
| AC mA | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 |  |
| AC A | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 |  |
| T1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 |  |
| T2 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 |  |
| Hz | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 |  |
| Harmonic Ratio | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 |  |
| % of mA | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 |

E. Range Definition for different functions

| Range | 0 | 1 | 2 | 3 | 4 | 5 |
| --- | --- | --- | --- | --- | --- | --- |
| Function |  |  |  |  |  |  |
| DC V | 999.9mV | 9.999V | 99.99V | 999.9V |  |  |
| DC µA | 400.0 | 4.000mA |  |  |  |  |
| DC mA | 40.00mA | 400.0mA |  |  |  |  |
| DC A | 4.000A | 10.00A |  |  |  |  |
| Ts | K/C | J/C | K/F | J/F |  |  |
| DIODE | 2.000V |  |  |  |  |  |
| OHM | 400.0 | 4.000k | 40.00k | 400.0k | 4.000M | 40.00M |
| Cx | 4.000F | 40.00F | 400.0F | 9999F |  |  |
| AC V | 999.9mV | 9.999V | 99.99V | 999.9V |  |  |
| AC µA | 400.0 A | 4.000mA |  |  |  |  |
| AC mA | 40.00 mA | 400.0 mA |  |  |  |  |
| AC A | 4.000A | 10.00A |  |  |  |  |
| T1 | K/C | J/C | K/F | J/F |  |  |
| T2 | K/C | J/C | K/F | J/F |  |  |
| Hz | 9.999Hz | 99.99Hz | 999.9Hz | 9.999kHz | 99.99kHz | 999.9k |
| Harmonic Ratio | 099.9% |  |  |  |  |  |
| % of mA | 999.9% (4~20mA) | 999.9% (0~20mA) |  |  |  |  |

Example:

| Format | start | FUNC. | Range | Sign | V5 | V4 | V3 | V2 | V1 | BAT | Stop |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| FUNC. | 1F | 20~2F | 20 | 26 | 21 | 29 | 26 | 28 | 29 | 20 | 1E |
| DCV | +196.89 mV |  |  |  |  |  |  |  |  |  |  |
| DCA | +196.89 A |  |  |  |  |  |  |  |  |  |  |
| DcmA | +19.689 mA |  |  |  |  |  |  |  |  |  |  |
| DCA | +1.9689 A |  |  |  |  |  |  |  |  |  |  |
| Ts | +1968.9 C (K type) |  |  |  |  |  |  |  |  |  |  |
| Diode | +1.9689 V |  |  |  |  |  |  |  |  |  |  |
| OHM | +196.89  |  |  |  |  |  |  |  |  |  |  |
| Cx | +1.9689 F |  |  |  |  |  |  |  |  |  |  |
| ACV | +196.89 mV |  |  |  |  |  |  |  |  |  |  |
| ACA | +196.89 A |  |  |  |  |  |  |  |  |  |  |
| AcmA | +19.689 mA |  |  |  |  |  |  |  |  |  |  |
| ACA | +1.9689 A |  |  |  |  |  |  |  |  |  |  |
| T1 | +1968.9 C (K type) |  |  |  |  |  |  |  |  |  |  |
| T2 | +1968.9 C (K type) |  |  |  |  |  |  |  |  |  |  |
| Hz | +1.9689 Hz |  |  |  |  |  |  |  |  |  |  |
| H% | +196.89 H% |  |  |  |  |  |  |  |  |  |  |
| %mA | +19.689 mA *1 |  |  |  |  |  |  |  |  |  |  |

Note :  1.   %mA = ( mA - 4 ) / 16 * 100.00%

= ( 19.698 - 4 ) / 16 * 100.00%

= 98.1125 %